## Slide 1
ELT / PCS POLARIMETRY
Work Update & Instrument Trade Study
Polarimetric Differential Imaging for ELT / PCS
[Your Name] | Polarimetry Lead, ELT PCS
June 2026
Notes: SPEAKER NOTES – Slide 1: Title
Good [morning/afternoon]. Today I am here to give you an update on my work over the past month as Polarimetry Lead for ELT PCS, with a particular focus on the instrument trade study I have been conducting.
The trade study is the main focus of today. I will walk through seven existing high-contrast polarimetric instruments, what worked, what didn't, and what that means for how we should design the PCS polarimetry sub-system.
This builds on all the papers and supervisor notes I have been reviewing.

## Slide 2
Agenda
What we will cover today
ELT PCS | Polarimetry Work Update | June 2026
2
1
Work completed this month
2
PCS polarimetry science context (brief recap)
3
Instrument trade study (main focus)
SPHERE, IRDIS, ZIMPOL, GPI, SCExAO, CHARIS, VAMPIRES
4
Key design lessons for PCS
5
ELT-specific challenges
6
Next steps
Notes: SLIDE ORDER:
1. Title | 2. Agenda | 3. Work Completed | 4. Science Context
5. Trade Study Intro | 6. SPHERE+IRDIS | 7. ZIMPOL | 8. GPI+SCExAO
9. CHARIS+VAMPIRES
10. KEY DISCUSSION POINTS (HWP, K-mirror, coatings)
11. POLARISATION ABERRATIONS (calibration + compensation)
12. BEAM-SPLITTING OPTIONS (Wollaston, cube, IFU, pol. grating)
13. Key Design Lessons | 14. ELT-Specific Challenges
15. Proposed Architecture | 16. Next Steps | 17. Conclusions

## Slide 3
Work Completed This Month
Past ~6 weeks: deliverables and activities
ELT PCS | Polarimetry Work Update | June 2026
3
Documents & Planning
WP3.4 Polarimetry Management Plan (10-section Word document, professional format)
WP3.3 section inserted into PCS project plan (inputs/outputs, activities, timescale)
Science cases documented: disk PDI, rocky planet reflected light, young planet thermal emission
Receivables/deliverables tables defined
21-task WBS + Gantt schedule compiled
Literature & Trade Study
14 papers reviewed (284-page merged PDF) + team review notes (PCS Updated.docx)
Instrument benchmarking: mapped technical performance of 7 instruments
SPHERE/IRDIS Mueller matrix calibration paper (van Holstein et al. 2020), key lessons
SPHERE/ZIMPOL high-resolution polarimetry (Schmid et al. 2018), kHz modulation lessons
Anche et al. 2023, ELT polarisation aberration analysis: diattenuation & retardance defocus
Trade study table compiled and added to WP3.4 document
Notes: SPEAKER NOTES – Slide 3: Work completed
On the left side, the planning work. I have created a full WP3.4 management plan following the same professional format as the WP3 document, with 10 sections covering science case, boundary conditions, calibration strategy, work breakdown structure, schedule, receivables and deliverables. I have also inserted a WP3.4 Polarimetry section into the main project plan document.
On the right, the research work. I reviewed 14 papers in the PCS literature, with particular focus on the calibration papers for SPHERE/IRDIS and SPHERE/ZIMPOL. I also worked through the Anche et al. 2023 paper on polarisation aberrations in next-generation extremely large telescopes, this is particularly important because ELT has a much faster focal ratio than VLT (F/0.9 vs F/4) which makes polarisation aberrations much worse.
The output of all that reading is the trade study table, which is now Section 7 of the WP3.4 document. That trade study is what I want to walk you through in detail today.

## Slide 4
PCS Polarimetry: Why Do We Need It?
Science cases and contrast requirements
ELT PCS | Polarimetry Work Update | June 2026
4
PDI: polarised flux only (Q, U) → a CONTRAST technique for detecting scattered/reflected light against the stellar halo
Polarimetry: also measures total intensity I → fractional polarisation q,u = Q/I, U/I → microphysical characterisation (dust scattering for disks; atmosphere/clouds/surface for planets)
Protoplanetary Disks
Scattered-light imaging with PDI; Qphi/Uphi decomposes Stokes Q,U in azimuthal coords
Contrast: 10⁻³–10⁻⁴ in Qphi
Spectral range: visible + NIR
Reveals grain properties, disk structure, spirals, gaps, planets shaping disks
SPHERE has >100 systems imaged
ELT: 5× better resolution + fainter/ more distant targets
Debris Disks
Scattered-light PDI, same Qphi/Uphi technique as protoplanetary disks
Older, more tenuous systems → fainter Qphi signal than protoplanetary disks
Spectral range: visible + NIR
Traces dust morphology: rings, clumps, warps — often shaped by unseen planets
GPI debris-disk survey: ~100 stars observed (Esposito et al. 2020)
Challenge: distinguishing a polarised planet signal from disk structure itself
Young Planets (Thermal)
Hot young giant planets emit NIR light (L = 10⁻³ to 10⁻⁵ L_star)
Intensity contrast: ~10⁻⁵ to 10⁻⁶
Spectral range: NIR
Polarisation fraction ≤ few % (atmospheric scattering)
PDI not the primary technique here but useful for companion disk detection
ELT gives access to lower-mass planets at smaller separations
Reflected Light
Gas giants: realistic near-term target; rocky M-dwarf planets = aspirational case
Intensity contrast: ~10⁻⁷ (gas giant) to 10⁻¹⁰ (rocky Earth-like)
Spectral range: visible
Polarised contrast: 10⁻⁸ for rocky planets (Stam 2008, 2004); polarimetry gains ~10⁻² since star is unpolarised
IP must be corrected to 10⁻³ level
Requires: dual-beam, upstream HWP, complete Mueller matrix calibration
Notes: SPEAKER NOTES – Slide 4: Science context
Just to recap the science motivation before we get into the trade study.
There are three science cases driving the polarimetry design. The first is protoplanetary disk imaging using PDI. This is the mature science case where SPHERE has already been enormously successful, and ELT will give us higher spatial resolution and access to fainter and more distant systems.
The second is young planet detection in thermal emission. Here polarimetry is useful for detecting circumplanetary material around young companions, but the primary technique is direct imaging.
The third, and most challenging, is rocky planets in reflected light. This is the unique PCS science case. The intensity contrast is 10⁻⁷ to 10⁻¹⁰. However, the key insight from Stam et al. is that the PLANET is polarised (scattered light from its atmosphere) while the STAR is essentially unpolarised. This gives polarimetry a factor of roughly a hundred to a thousand advantage over direct intensity imaging.
But to exploit this, we need to correct instrumental polarisation to the 10⁻³ level. This is the fundamental requirement that drives almost every design decision for PCS polarimetry.

## Slide 5
Instrument Trade Study
Which instruments? What did we assess?
ELT PCS | Polarimetry Work Update | June 2026
5
Seven instruments assessed across three generations of high-contrast polarimeters
SPHERE
(General)
IRDIS
(SPHERE/VLT)
ZIMPOL
(SPHERE/VLT)
GPI
(Gemini South)
SCExAO
(Subaru)
CHARIS
(Subaru/SCExAO)
VAMPIRES
(Subaru/SCExAO)
Assessment framework for each instrument:
✅ What worked?
❌ What didn't work?
➡ Lessons for PCS design
Notes: SPEAKER NOTES – Slide 5: Trade study intro
We assessed seven instruments. They fall into three groups:
VLT/SPHERE instruments are the most directly relevant to PCS because SPHERE is the current state of the art for high-contrast polarimetry and shares many design elements. We look at both the NIR dual-beam channel (IRDIS) and the visible fast-modulation channel (ZIMPOL).
Gemini: GPI, the Gemini Planet Imager is interesting mainly because of its IFS-based approach to polarimetry, and also because it's at Cassegrain focus rather than Nasmyth, which makes IP calibration simpler.
Subaru/SCExAO: three instruments here, SCExAO itself is the platform, CHARIS the NIR IFS, and VAMPIRES the visible fast-modulation polarimeter. VAMPIRES is particularly instructive because it uses FLC modulation and comparing it to ZIMPOL helps us decide what modulation strategy is right for PCS.
For each instrument I asked: What worked? What didn't? And what does that mean for PCS design?

## Slide 6
Trade Study: SPHERE (Overall) & SPHERE/IRDIS
NIR dual-beam polarimetry | van Holstein et al. 2020, 2021, 2023
ELT PCS | Polarimetry Work Update | June 2026
6
SPHERE — General System Performance
✅ Worked
• PDI (Qphi imaging): gold standard; >100 systems imaged
• Coronagraphy + PDI → 10⁻³–10⁻⁴ contrast in Qphi
• Uphi diagnostic: ≈ 0 when well-calibrated; key quality metric
• Dual-beam effectively removes AO speckle residuals
❌ Problems
• Nasmyth M3 mirror: 1.5–3.5% telescope IP (wavelength dependent)
• Large calibration overheads (~1 hr per target + standards)
• Bright stars only: R < 11 for ExAO performance
• IP residual after calibration: ~0.5% (limits inner-working-angle)
➡ Lessons for PCS
• PDI is THE technique for disk science, include from Day 1
• ELT Nasmyth M4+M5 will introduce similar IP, HWP before fold mirrors is essential
• Mueller matrix model of ELT + PCS must be built before first light
• Uphi images: always report as calibration quality diagnostic
SPHERE/IRDIS — NIR Dual-Beam Polarimeter (J, H, K bands)
✅ Worked
• Full Mueller matrix calibration model, first for an AO instrument
• IRDAP pipeline: publicly available, validated on 100s of datasets
• Absolute polarimetric accuracy: <0.1% (relative: <1%)
• HWP cycling + dual-beam removes first-order IP
• First detection of circumsubstellar disk around DH Tau B
❌ Problems
• Derotator can drop to 5% (H) / 7% (K_s) min. efficiency in worst case; typical observing strategy is much better
• Derotator: |m₃₄| > 0.5 crosstalk (linear↔circular) at some angles
• Derotator at 45°: destroys polarimetric efficiency entirely
• M3 vs M4 coatings not identical despite design intent; drifts further with coating ageing
• Pol. beamsplitter = polarizers BEHIND a regular beamsplitter → factor 2 photon loss
➡ Lessons for PCS
• HWP retardance must be uniform close to 180° in ALL science bands
• Derotator (K-mirror): keep away from 45°, optimise coatings, model every state
• ELT F/0.9 polarisation aberrations (not simple beam shifts) scale as 1/f² → ~20× worse than SPHERE F/4
• Polarisation-aberration correction must be built into PCS data pipeline from Day 1
• IRDAP-style public calibration pipeline is essential
Notes: SPEAKER NOTES – Slide 6: SPHERE + IRDIS
Starting with SPHERE which is the most directly relevant comparison instrument for PCS.
At the top level, SPHERE PDI has been transformatively successful. The Qphi imaging technique, where you measure polarisation in azimuthal coordinates around the star, is now the standard approach for disk imaging. SPHERE has imaged over 100 protoplanetary disk systems.
But the main limitation is the Nasmyth configuration. The M3 fold mirror introduces 1.5 to 3.5% telescope polarisation depending on wavelength and pointing. This needs to be calibrated away, and that calibration takes time.
Moving to IRDIS specifically, the NIR dual-beam channel. This is where the real technical work happened. The team around Rob van Holstein built the first complete Mueller matrix model of an adaptive optics instrument, and this became the IRDAP pipeline that the community now uses.
The big problems for IRDIS: first, the HWP has terrible retardance performance in H band only about 5% of the polarimetric signal survives in the worst case. This is a manufacturing defect essentially. For PCS the specification on the HWP must be much tighter.
Second, the derotator. This is a K-mirror type that rotates to compensate for field rotation in alt-az telescopes. The problem is that a K-mirror at 45 degrees completely destroys polarimetric efficiency. And at other angles it introduces large crosstalk between linear and circular polarisation. This must be modelled per exposure.
Third, and this is a relatively new discovery from van Holstein 2023, there are polarisation beam shifts at every inclined mirror. The Goos-Hänchen and Imbert-Federov shifts cause the orthogonally-polarised beams to land at slightly different positions on the detector. For SPHERE this limits the contrast gain from PDI to about a factor of 350. For ELT at F/0.9, these effects will be much larger because beam shifts scale inversely with the f-number.

## Slide 7
Trade Study: SPHERE/ZIMPOL
Visible fast-modulation polarimetry | Schmid et al. 2018
ELT PCS | Polarimetry Work Update | June 2026
7
SPHERE/ZIMPOL — Zürich IMaging POLarimeter (500–900 nm, 3.6 mas/pixel)
✅ What Worked
• kHz EMCCD modulation: I⊥ and I∥ measured on SAME PIXELS removes all atmospheric/speckle temporal variation
• Highest spatial resolution optical polarimetry on 8m: 3.6 mas pixels; PSF 20–22 mas
• Polarimetric fidelity over >3000:1 intensity dynamic range (clean Q/U images throughout)
• Multiple coronagraph options (CLC, 4QPM) simultaneously available
• Field/polarisation angle offsets allow disentangling instrumental from sky signals
• Detected circumstellar flux at Cpol > 10⁻⁴ within 0.2 arcsec
• Several HWPs steer polarisation through the optical train — some nice-to-have, some (e.g. before the derotator) likely crucial
• One HWP is used together with the FLC in a 2-stage modulation scheme — an elegant combination
❌ Problems
• Limited AO performance in the visible is ZIMPOL's single biggest limiting factor
• VLT M3 Nasmyth mirror (Al-coated): ~4% IP in 500–900 nm range
• Derotator: ~3% IP + |m₃₄| > 0.5 wavelength-dependent crosstalk
• No pupil-stabilised polarimetry → cannot combine PDI + ADI
• Visible only (500–900 nm), blocking filter inadequate beyond this band: unsuitable for young hot planets in thermal emission
• EMCCD pixel faults from charge-shifting require extra calibration steps
• ~15% throughput loss from polarimetric optics
➡ Lessons for PCS
• Fast modulation eliminates speckle noise in differential images, very powerful principle
• Faster modulation is always better in principle, but kHz needs a special detector that may add more complexity/low-light issues than the gain justifies
• A true dual-beam system (IRDIS-style) could work just as well — ZIMPOL's 2 detectors were never designed identically, so it's effectively a 2× single-beam system
• PCS MUST cover NIR as well as visible, IRDIS-style NIR channel needed for young planet thermal emission
• Two-beam same-pixel design is the ideal architecture for polarimetric differential imaging
• ELT Nasmyth IP (M4+M5) mirrors SPHERE problem, HWP upstream of fold mirrors is the solution
• ADI + PDI combination is needed, requires pupil-stabilised polarimetry (not offered by ZIMPOL)
Notes: SPEAKER NOTES – Slide 7: ZIMPOL
ZIMPOL is the visible channel of SPHERE and is really the most innovative polarimeter currently operating on a large telescope.
The key innovation is kHz EMCCD modulation. The detector alternates between measuring the I-perp and I-parallel polarisation states at kilohertz frequency. Because the two states are measured on the same physical pixels, any flat-field effects, gain variations, or speckle patterns cancel almost perfectly. This is why ZIMPOL achieves excellent polarimetric fidelity, the Q/U images are clean even with a 3000:1 intensity dynamic range.
The spatial resolution is spectacular: 3.6 mas pixels with PSF around 20 to 22 mas in good conditions. That's about 5 times better than IRDIS.
But the problems: same M3 Nasmyth issue as SPHERE generally about 4% IP at visible wavelengths. The derotator is even worse in some respects, it is a 3-mirror system with two 55-degree mirrors, and the crosstalk between linear and circular polarisation can exceed 50%.
The biggest operational limitation is that ZIMPOL doesn't support pupil-stabilised polarimetry. This means you can't combine PDI and ADI (Angular Differential Imaging), which is what gives the best possible speckle suppression. IRDIS can do pupil-stabilised mode, which is why it's often used for companion searches.
Now, the most important lesson for PCS: after talking with my supervisor, he notes explicitly that the kHz FLC modulation, which is what VAMPIRES at Subaru also uses, is NOT necessary. The key is to use an HWP that rotates to different angles to modulate between Q and -Q. Doing this at 2Hz rather than kHz is sufficient because the atmospheric speckles are suppressed by the dual-beam design, not by the modulation speed. This simplifies the instrument enormously.
PCS must cover NIR as well. Young planets emit thermally in J/H/K bands, and we need to be able to detect them. So, we need something like ZIMPOL for visible AND something like IRDIS for NIR.

## Slide 8
Trade Study: GPI & SCExAO
Gemini Planet Imager & Subaru Coronagraphic Extreme AO
ELT PCS | Polarimetry Work Update | June 2026
8
GPI (Gemini South) | Lenslet IFS + Wollaston | J, H, K_s bands | Cassegrain
✅ Worked
• IFS spectropolarimetry simultaneously in J/H/K_s
• Cassegrain mount: NO fold-mirror IP, greatly simplifies calibration
• GPIES: >500 stars surveyed; disks + companions detected
• Wollaston prism in pupil plane: high polarisation purity
• Public pipeline (GPI-DRP) community-validated
❌ Problems
• Lenslet limits spatial resolution vs. dedicated imager
• ~20% throughput loss from IFS optics
• No visible wavelength capability (> 900 nm only)
• Complex IP calibration through many surfaces; M1 polarisation-astigmatism effects visible in GPI's polarised PSF
➡ Lessons
• Cassegrain position simplifies IP vs. Nasmyth, ELT PCS must compensate via HWP
• Wollaston prism in pupil plane: validated for PCS NIR beam splitting
• GPI offers IFS spectroscopy OR IFS polarimetry, never both at once
• Dual-beam works perfectly when both split states are sampled identically by a microlens array (MLA) before dispersion
• IFS spectropolarimetry is an opportunity GPI didn't combine — PCS could
• Public pipeline from Day 1 (GPI-DRP model)
SCExAO (Subaru) | Modular ExAO platform | hosts CHARIS, VAMPIRES, MEC
✅ Worked
• Modular design: rapid upgrade path (CHARIS, VAMPIRES, MEC)
• Best NIR Strehl on 8m (>90% H band in good conditions)
• LLOWFS demonstrated: post-coronagraphic wavefront control
• Speckle nulling + electric field conjugation validated
❌ Problems
• Nasmyth platform: silver-mirror IP ~2–3% NIR
• Linear polarisation is no longer an eigenvector of the full mirror train everywhere — adds real complexity
• AO performance drops rapidly for R > 11 guide stars
➡ Lessons
• Main lesson: SCExAO is a vital platform for experimentation, but PCS needs FAR less complexity
• LLOWFS for coronagraph stabilisation is critical at ELT scale
• IP from Nasmyth is unavoidable, upstream HWP is essential
• Speckle nulling at ELT will need more DM actuators
Notes: SPEAKER NOTES – Slide 8: GPI and SCExAO
GPI is interesting for PCS because of two things: first, the Cassegrain mounting, and second, the IFS approach.
Being at Cassegrain means GPI doesn't have a Nasmyth fold mirror, so there's no IP from that source. This makes the calibration much simpler. ELT/PCS will almost certainly be at Nasmyth, this is already decided by the telescope design, so we have to deal with M4 and M5 introducing IP. The lesson is that we need an HWP upstream of those mirrors.
The IFS approach, lenslet array plus Wollaston prism, gives you spectropolarimetry: you measure both polarisation and the spectrum simultaneously. This is scientifically very powerful, especially for characterising companion atmospheres. But for disk imaging it's not ideal because the spatial resolution suffers compared to a dedicated imager. GPI-DRP, the public pipeline, has been widely used by the community. I will speak more on the Wollaston prims a bit later because I think this is very pertinent to what we want to do in the end.
SCExAO at Subaru is valuable as a concept study for the modular approach. It's a platform that hosts multiple science instruments simultaneously. This gives it great versatility but also creates calibration complexity because all these instruments share the beam. The main achievement for SCExAO relevant to polarimetry is the LLOWFS, the Lyot Low-Order Wavefront Sensor, which provides real-time monitoring of the wavefront quality behind the coronagraph. At ELT, where the inner working angle will be much tighter, this kind of sensor will be even more critical.

## Slide 9
Trade Study: CHARIS & VAMPIRES
Subaru/SCExAO NIR IFS and visible dual-beam polarimeter
ELT PCS | Polarimetry Work Update | June 2026
9
CHARIS (Subaru/SCExAO) | NIR lenslet IFS | J+H+K simultaneous | 1.1–2.4 μm
✅ Worked
• J/H/K band simultaneous coverage (1.1–2.4 μm)
• R~18 (broadband) to R~70 (narrowband), companion spectra
• Spectropolarimetry mode with Wollaston demonstrated
• Directly imaged planet characterisation (HR 8799 etc.)
❌ Problems
• No dedicated fast modulation → speckle noise in pol. images
• Wollaston sits in FRONT of the lenslet (MLA) — fixed pattern + mismatched sampling can mimic polarisation signals
• Complex cross-channel calibration across 1.1–2.4 μm
• ~20–30% throughput penalty from lenslet + prism optics
• Very limited FOV (1.1″ × 1.4″)
➡ Lessons
• Dedicated polarimetric channel is better than IFS alone for disk imaging
• IFS mode (CHARIS-style) ideal for companion characterisation as SECONDARY mode
• PCS design: dedicated polarimetric imager for disks + IFS option for companions
• Spectropolarimetry of companions is a powerful PCS capability to include
VAMPIRES (Subaru/SCExAO) | Visible dual-beam fast-mod. | 600–800 nm | FLC modulation
✅ Worked
• Two-camera simultaneous I⊥ + I∥: eliminates temporal atmospheric noise
• FLC fast modulation (~100 Hz) suppresses speckle variations
• Demonstrated on AGB stars, stellar winds, protoplanetary disks
• Polarisation contrast ratio >1000:1 achieved routinely
• Aperture masking mode: interferometric resolution
❌ Problems
• FLC modulator: temperature sensitivity, long-term drift, maintenance overhead
• FLC chromatic retardance → complex wavelength-dependent calibration
• Nasmyth IP: ~2–3% in visible
• FLC + derotator: complex time-varying Mueller matrix
• 2 QWPs needed just to get polarisation through the periscope — 3 modulation stages total, all super complex
➡ Lessons
• Two-camera SIMULTANEOUS design is the single most important architectural choice
• Modulate as fast as the detector allows: ~100 Hz FLC is still interesting, a rapidly spinning HWP may be too — but true dual-beam needn't beat the speckle lifetime
• Visible polarimetry (600–900 nm) scientifically important for rocky planet reflected light
• HWP is more achromatic than FLC → better for the broad wavelength range of PCS
• Dual-beam design removes flat-field uncertainty: must be central to PCS architecture
Notes: SPEAKER NOTES – Slide 9: CHARIS and VAMPIRES
CHARIS at Subaru is the NIR IFS counterpart to SPHERE/IRDIS, but covers a much wider wavelength range simultaneously, J, H, and K bands together. The spectropolarimetry mode is scientifically very powerful: you get both the spectrum and the polarisation of a companion in one observation.
But for disk imaging specifically, the lenslet sampling reduces the effective spatial resolution. And the cross-channel calibration for polarimetry across 1.3 magnitudes of wavelength range is genuinely difficult. The lesson for PCS is that we should have a dedicated polarimetric channel for disk science, and treat IFS spectropolarimetry as a secondary mode for companion characterisation.
VAMPIRES is the most instructive comparison for the modulation strategy decision. VAMPIRES uses FLC, Ferroelectric Liquid Crystal, modulators that switch between polarisation states at ~100 Hz. This effectively suppresses all atmospheric speckle variations.
But in practice, FLC modulators are fragile: they're temperature sensitive, they drift over time, and they need frequent calibration. They also have chromatic retardance that varies across the wavelength range.
Again, talking with my supervisor he explicitly address this. The message is: FLC fast modulation is not necessary for PCS. The key insight is that the dual-beam simultaneous design already removes the temporal atmospheric noise, you just need to modulate between Q and -Q, and you can do that with an HWP rotation at 2 Hz rather than an FLC at 100 Hz. In any case, a HWP is more achromatic, more stable, and much simpler to maintain.
The two-camera simultaneous design is the essential element. That's what VAMPIRES gets right. The modulator just needs to switch the sign of Q, and any reasonably fast HWP does that.

## Slide 10
Key Design Discussion Points
Critical decisions for PCS polarimetry architecture
ELT PCS | Polarimetry Work Update | June 2026
10
The following three topics represent the most critical design decisions — flagged for immediate action:
1
HWP as the Primary Modulator
• Rotate the polarisation coordinate system using the HWP — this is the core modulation strategy
• HWP at 0°, 22.5°, 45°, 67.5° → measures +Q, +U, -Q, -U (complete Stokes Q and U)
• Placement now uncertain: PCS likely sits behind facility M6 at folded Nasmyth — open question is modulator before M6, not before M4/M5
• Must be achromatic: retardance 180° ± 2° across full science wavelength range (J through K)
• Rotation speed: as fast as possible — kHz FLC may not be necessary; exact requirement depends on AO/dark-hole performance
• HWP upstream means telescope IP is modulated away and calibrated to first order automatically
2
K-Mirror: Keep Away from 45°, Calibrate Efficiency Losses
• K-mirror is a 3-flat-mirror image derotator used to compensate field rotation in alt-az telescopes
• The K-mirror acts like a QWP, not a HWP — not all angles are problematic; if aligned S/P with the other mirrors, there's no issue
• Must define operational constraint: avoid K-mirror angle = 45° during science observations
• AT ALL angles: K-mirror introduces wavelength-dependent crosstalk (|m₃₄| can exceed 0.5)
• Calibrate efficiency losses using dedicated calibration sequences at each K-mirror angle
• Log K-mirror angle in ALL science headers — essential for post-processing calibration
• Coordinate with telescope/instrument team to confirm operational avoidance strategy is feasible
3
Optimise Mirror Coatings for Polarisation, Not Just Reflectivity
• Standard astronomical coating optimisation: maximise reflectivity = minimise diattenuation
• For polarimetry: we ALSO need to minimise retardance (ideally = 0° or 180°, not 90°)
• These goals conflict: most reflective coatings (Al, Ag) have strong wavelength-dependent retardance
• Approach: for each mirror, jointly optimise diattenuation AND retardance over our science bands
• Challenge: K-mirror and Nasmyth mirrors (M4, M5) cover J/H/K — hard to optimise simultaneously
• Anche et al. 2023 showed: coating optimisation can reduce ELT diattenuation below 1% while maintaining high reflectivity
• This is a design input for the ELT mirror specification — we must engage with the ESO optics team early
Van Holstein et al. (2019)
Notes: SPEAKER NOTES — Key Discussion Points
These are the three most critical near-term design decisions that I want to flag with the team.
Number 1: the HWP as our primary modulator. The principle is simple: we rotate a half-wave plate to four positions: 0, 22.5, 45, and 67.5 degrees and this rotates the polarisation coordinate system into and out of our detector's measuring axes. At these four positions we measure +Q, +U, -Q, and -U. From the differences we reconstruct Q and U. The critical constraint is that the HWP must be upstream, before any Nasmyth mirrors, so that telescope IP is modulated away.
Number 2: the K-mirror. I want to be explicit: the K-mirror is the image derotator, they're the same component. At exactly 45 degrees, the K-mirror geometry acts like a HWP itself, which means it completely negates our HWP modulation and our polarimetric efficiency goes to zero. We need a firm operational constraint: never observe science targets when the K-mirror is at 45 degrees. At all other angles there are efficiency losses that we need to calibrate.
Number 3: coatings. Most coating specifications optimise for reflectivity, which is what astronomers want for sensitivity. But for polarimetry we also need to specify the retardance. If a mirror has a retardance close to 90 degrees, it acts like a quarter-wave plate and scrambles linear into circular polarisation which is very bad for us. The Anche et al. paper showed it IS possible to jointly optimise for both reflectivity and low diattenuation. We need to get this into the ELT mirror coating specification early, before those specifications are frozen.

## Slide 11
Polarisation Aberrations: The ELT Design Challenge !!
Impact on coronagraphic performance and calibration strategy — Anche et al. 2023
ELT PCS | Polarimetry Work Update | June 2026
11
⚠ What are polarisation aberrations?
• When a beam converges/diverges at a mirror, different rays hit at different incidence angles
• The Fresnel equations give DIFFERENT retardance/diattenuation for each ray across the beam
• This creates a spatially-varying polarisation pattern across the pupil (Jones pupil aberrations)
• Result: PSF is not the same for I⊥ and I∥ → residual signal in Q/U images that mimics science signal
• Beam shift = polarisation tip/tilt → PSF splits, gradient-like structures appear in Q, U
• M1's effect = polarisation astigmatism (diattenuation + retardance) → PSF rings/wings, speckles become polarised; rings azimuthally polarised like a disk, speckles randomly polarised but only slowly average out
• Effect scales as 1/f² → ELT at F/0.9 is ~20× worse than VLT at F/4 (segmentation adds further structure)
⚠ Impact on coronagraphic performance
• Anche et al. 2023: polarisation aberrations limit coronagraphic contrast to ~10⁻⁵ (H/K bands)
• The stellar residuals are mainly DEFOCUS in shape, not speckle-like → hard to subtract
• This affects BOTH coronagraphy (total intensity) AND polarimetry (Q/U images)
• For polarimetry: the azimuthal polarisation structure mimics scattered light from a circumstellar disk
• Disk science requires we understand and subtract this at the 10⁻³ level or better
❓ Can we calibrate for the systematic effects?
• YES — in principle. Full Mueller matrix model includes polarisation aberrations as spatial terms
• Calibration approach: measure PSF in Q and U using unpolarised ("boring") reference stars → model the pupil-plane polarisation map
• Remaining question: DOWN TO WHAT LEVEL can we calibrate?
◦ At 10⁻³ IP: likely achievable with current IRDAP-type approaches adapted for ELT
◦ At 10⁻⁴: polarisation aberrations become dominant; new calibration strategy needed
◦ At 10⁻⁵: approaching the photon-noise floor; spatially resolved pupil calibration required
• Time-varying component: mirror coating aging changes polarisation properties over years
• Must define: how often do we re-calibrate? What is the stability timescale?
!! Can we design a compensating component in the pupil plane?
• Concept: a spatially-varying retarder (patterned waveplate) in the pupil plane
• Designed to cancel the retardance aberration map from M1–M5
• Similar to a "pupil-plane coronagraphic apodiser" but for polarisation
• Liquid crystal (LC) technology can create spatially patterned waveplates
• Challenge: must compensate BROADBAND (chromatic LC) AND across the optic — we currently only coat with uniform retardance at CLJ
• This is a potential Phase A design study: quantify benefit vs. complexity cost
• May significantly relax calibration requirements → worth serious investigation
Notes: SPEAKER NOTES — Polarisation Aberrations
This slide addresses what I consider the most serious technical risk for ELT/PCS polarimetry, polarisation aberrations.
The problem: when a beam converges onto a mirror at a fast focal ratio like F/0.9, different rays in the beam hit the mirror at different angles of incidence. Because the Fresnel equations give different amplitude and phase shifts for different incidence angles, the pupil ends up with a spatially-varying polarisation pattern. This creates what's called a Jones pupil aberration. The result is that your PSF in the Q image looks different from your PSF in the U image, you get a residual polarised signal even from an unpolarised star.
The most dangerous term is "retardance defocus", this creates an azimuthal ring of polarised light around the PSF, which for our disk science looks exactly like scattered light from a circumstellar disk. The Anche et al. 2023 paper quantified this for ELT and found it limits coronagraphic contrast to about 10 to the minus 5.
For our POLARIMETRY this is a key question, well, two questions really:
First: can we calibrate these systematic effects? In principle yes. We measure the polarisation aberration map using unpolarised standard stars and build a model that we subtract. The question is: how precisely? At 10^-3 (what we need for disk science) it should be achievable. At 10^-4 it gets harder. At 10^-5 (for rocky planet science) we probably need something more.
Second: should we design a physical compensating element? This is really the exciting option that I want the team to consider. We could put a patterned waveplate, a spatially-varying retarder, in the pupil plane, specifically designed to counteract the polarisation aberration pattern from the ELT optics. Liquid crystal technology can make these. The challenge is making it work broadband. But if it works, it could completely eliminate or dramatically reduce polarisation aberrations, which would relax all our other calibration requirements.
This is something I'd propose as a dedicated Phase A study: quantify the benefit of a compensating optic versus the complexity cost.

## Slide 12
Beam-Splitting Architecture Options
How we split the two polarisation states determines the whole instrument design
ELT PCS | Polarimetry Work Update | June 2026
12
Choice of beam splitter is a fundamental architecture decision — it affects detector layout, calibration, spectral coverage, and throughput:
A
Wollaston Prism in PUPIL plane
• Separates two beams in PUPIL, each beam is a full image of the source on half the detector
• Both polarisation states simultaneously on same detector: excellent flat-field cancellation
• Clean angular separation: I⊥ and I∥ beams land side-by-side on detector
• Calibration: beam shear must be measured and corrected (the GH/IF beam shifts)
• Used by: SPHERE/IRDIS, GPI (both validated)
• → Preferred choice for dedicated NIR polarimetric imaging channel
B
Cube Beam Splitter in IMAGE plane
• Transmits one polarisation, reflects other at 90°, two separate detector arms
• Image-plane split: better PSF sampling but requires two detectors in precise alignment
• Higher throughput for individual channels (no grating loss)
• Chromatic beamsplitter cubes have wavelength-dependent splitting ratio → calibration overhead
• IRDIS uses regular polarizers behind a regular beam-splitter, not a polarizing cube
• → Valid option; simpler optics but requires detector co-registration
C
GPI approach: IFU + Wollaston for spectropolarimetry
• Lenslet IFS where grating is replaced by / supplemented with Wollaston prism
• Each lenslet gives I⊥ and I∥, you get BOTH spectral AND polarisation info
• Can be done with POLARISATION GRATINGS (grating that simultaneously disperses AND separates pols)
◦ Polarisation gratings: single element replaces grating + Wollaston → more compact
◦ Sampled identically in both pol. states → minimal systematic errors
• → Enables full spectropolarimetric IFU for companion science; cost: lower throughput
D
Polarisation Grating + Quarter-Wave Plate
• Polarisation grating (liquid crystal) splits left- and right-circular polarisations
• Combined with a QWP ahead of it: converts linear Q/U to circular → grating then separates them
• Much wider acceptance angle than Wollaston (ELT's fast beam itself isn't the deciding factor here)
• Works across a broad wavelength range if well-designed; less sensitive to input beam angle than Wollaston (important for ELT pupil quality)
• → Attractive alternative to Wollaston; worth evaluating for PCS; less heritage but growing use
Architecture depends on science priority: disk imaging (A or D) | companion spectropolarimetry (C) | survey efficiency (B). Options can be combined.
Notes: SPEAKER NOTES — Beam-Splitting Architecture Options
This slide is about the four main ways we can split the two polarisation states in PCS. This is a fundamental design choice because it determines the detector layout, the calibration approach, and what science modes are available.
Option A: Wollaston prism in the pupil plane. This is the traditional high-contrast polarimetry approach, used in both SPHERE/IRDIS and GPI. The Wollaston takes the pupil and deflects the two orthogonal polarisation states in opposite directions. You get two images of the source side by side on the same detector, both measured simultaneously. The flat-fielding errors are minimal. This is the validated approach for dedicated disk imaging.
Option B: cube beam splitter in the image plane. This reflects one polarisation and transmits the other, creating two separate optical paths that go to two separate detectors. IRDIS actually does something like this as well. The advantage is potentially better PSF sampling and higher throughput for each individual detector. The challenge is precise co-registration of the two detector arms.
Option C: the GPI-inspired IFU approach. Instead of using a grating to disperse the spectrum in an IFS, you replace or supplement it with a Wollaston or polarisation grating. Each lenslet then gives you two polarisation channels rather than a spectrum. Or, and this is the really interesting possibility, use a polarisation grating that simultaneously disperses AND separates polarisation states. This would give you a full spectropolarimetric IFU from a single element. The trade-off is throughput, but with a low-noise detector this could be very powerful.
Option D: polarisation grating plus quarter-wave plate. This is the alternative to Wollaston that's gaining use in modern instruments. The QWP converts linear to circular polarisation, and then the polarisation grating splits left and right circular. The big advantage is that polarisation gratings work at a wider range of input beam angles than Wollaston prisms, which matters a lot for ELT's fast F/0.9 beam. They can also be manufactured to work across a broad wavelength range.
The bottom line: the right choice depends on our primary science case. For disk imaging, A or D is preferred. For companion spectropolarimetry, C is very attractive. We may want to include more than one option in different instrument modes.
This is something I would like the polarimetry team's input on, particularly whether the polarisation grating option (D) is mature enough for PCS, and whether the spectropolarimetric IFU (C) is worth the complexity.

## Slide 13
Key Design Lessons for PCS Polarimetry
Synthesised from 7-instrument trade study
ELT PCS | Polarimetry Work Update | June 2026
13
1
HWP modulator — NOT FLC
HWP rotation as fast as possible: simpler, more achromatic, more stable than FLC; kHz may not be needed depending on AO/dark-hole performance. Upstream of M4/M5 is unlikely — open question is before M6. Retardance 180° ±2° across target band. Separate HWPs for visible and NIR.
2
Dual-beam simultaneous design
Measure I⊥ and I∥ simultaneously, sampled as identically as possible — removes temporal atmospheric and speckle noise. Wollaston prism in pupil plane preferred. Both pol. states on same detector.
3
Mueller matrix calibration model
Full Mueller matrix model of ELT+PCS from M1 to detector — Day 1. Must include full time variability of M4,5-rotation-M6, wavelength dependence, and long-term coating ageing. Calibrate with internal source + unpolarised standard stars. Build IRDAP-style public pipeline.
4
Derotator (K-mirror): model every state
K-mirror introduces strong wavelength-dependent cross-talk. Avoid K-mirror = 45° (efficiency goes to zero). Optimise coatings to minimise retardance. Log all angles per exposure.
5
Polarisation-aberration correction
ELT F/0.9 induces polarisation aberrations (not simple beam shifts) that worsen as 1/f² — ~20× worse than SPHERE F/4. Correction must be built into the calibration pipeline from Day 1.
6
NIR + Visible dual channel
NIR (J/H/K) for young planet thermal emission and disk science. Visible (600–1000 nm) for rocky planets in reflected light. Separate pol. optics per channel mean they are never used simultaneously. HWP achromaticity limits: use separate optimised HWPs per channel.
7
Modulator BEFORE AO — a common question
Modulator at intermediate focus is not going to happen. The real open question is whether it can go before M6 — modulating as fast as possible there would not interfere with AO at kHz. Validated principle: de Juan Ovelar et al. 2014 + SPHERE heritage.
Notes: SPEAKER NOTES — Key Design Lessons (UPDATED: added Lesson 7)
1. HWP modulator: prefer HWP rotation at 2 Hz over FLC. Simpler, less temperature-sensitive, more achromatic.
2. Dual-beam: both pol states simultaneously. Wollaston in pupil plane.
3. Mueller matrix calibration: full optical model, IRDAP-style pipeline, public release.
4. K-mirror: avoid 45 degrees. Model at all angles. Log in headers.
5. Beam shifts: GH+IF shifts scale as 1/f. ELT F/0.9 is 5x worse than SPHERE.
6. NIR + Visible split: need separate optimised HWPs per channel — single HWP cannot cover both J-K and R-I bands with acceptable retardance accuracy.
7. LESSON 7 (new, commonly asked): Can you put the modulator BEFORE the AO system?
Answer: YES, and it's actually the BEST place. The ELT intermediate focus (IF) sits between M3 and M4. At this point the beam has only bounced off M1, M2, M3 — all rotationally symmetric mirrors that contribute zero IP. M4 (the adaptive mirror) and M5 are the ones that generate IP. So placing the HWP at the IF means it modulates the signal BEFORE the IP is introduced, which is mathematically ideal.
People worry that spinning the HWP at 2 Hz will affect the AO. It doesn't — AO runs at ~1000 Hz and the HWP position changes every 0.5 seconds. No AO loop would see this as anything but a static phase offset.
de Juan Ovelar et al. 2014 (A&A 562, A8) formally simulated this for the E-ELT and showed the switch technique (HWP at IF) brings linear IP below requirements.
SPHERE: the HWP is placed before the SAXO AO module. This is the standard design.

## Slide 14
ELT-Specific Challenges
Why ELT polarimetry is harder than at 8m telescopes — Anche et al. 2023
ELT PCS | Polarimetry Work Update | June 2026
14
🔴 Fast focal ratio F/0.9 vs. F/4
• Polarisation aberrations (not simple beam shifts) scale as 1/f² → ~20× worse than SPHERE
• Polarisation aberrations (Anche 2023): retardance defocus and tilt much larger
• Creates azimuthal polarisation structure in PSF that mimics disk signal
• Peak contrast limited to ~10⁻⁵ at H/K bands from polarisation aberrations alone
🔴 M4 + M5 + M6 at Nasmyth platform
• Both mirrors introduce IP; M4 and M5 angles change with pointing
• IP from these must be modelled and subtracted to reach 10⁻³ level
• Silver-coated mirrors have wavelength-dependent retardance → complex IP spectrum
• Mitigation: place HWP before M4 (as far upstream as possible in PCS optical path)
🔴 Segmented primary M1
• Segment-to-segment coating variation → non-uniform pupil polarisation
• Missing segments, edge diffraction create extra polarisation structure
• Ashcraft 2025: simulations include segment errors and coating non-uniformity
• May require spatial modelling of pupil polarisation, complex calibration
🔴 Scale & AO complexity
• ELT: 39m primary, >6000 DM actuators, ~10× more sub-apertures than SPHERE
• AO residuals at ELT are different in character from 8m → need to validate PDI model
• NCPA (non-common-path aberrations) harder to control at ELT scale
• Much larger data volumes and calibration overheads
Notes: SPEAKER NOTES – Slide 11: ELT-specific challenges
This slide is important because it explains why we can't just copy SPHERE. The ELT has several features that make polarimetry substantially harder.
First: the focal ratio. ELT at the Nasmyth focus has F/0.9. SPHERE/VLT works at F/4. Polarisation beam shifts, those Goos-Hänchen and Imbert-Federov effects I mentioned, scale inversely with f-number. So, we are looking at roughly 4 to 5 times worse beam shift effects than at SPHERE. And Anche et al. 2023 showed that the polarisation aberrations, specifically retardance defocus and tilt, are much larger for ELT than for 8m telescopes. These aberrations create an azimuthal polarisation structure in the PSF that can mimic the signal from a circumstellar disk. That's a serious problem.
Second: the Nasmyth platform. ELT has M4 and M5 on the Nasmyth platform. Both mirrors change their incidence angle with telescope pointing, so the instrumental polarisation is not constant across the sky. We need a complete pointing-dependent model.
Third: the segmented primary. The 39m primary has nearly 800 hexagonal segments, each with its own coating. If those coatings age at different rates, the pupil becomes non-uniformly polarised. Anche et al. are planning to simulate this in their follow-up paper.
Fourth: the sheer scale. Six thousand DM actuators means a very different AO residual error structure, and much larger data volumes requiring automated calibration pipelines.
The bottom line from Anche et al.: polarisation aberrations limit the achievable contrast to about 10⁻⁵ between H and K bands even with perfect coronagraphy. To go beyond that for rocky planet science, we need either compensation optics or spatial calibration of the PSF structure.

## Slide 15
Proposed PCS Polarimetry Architecture
Design choices informed by trade study
ELT PCS | Polarimetry Work Update | June 2026
15
Based on lessons from all 7 instruments, the following architecture is proposed for PCS:
HWP Modulator
Located after M5/M6 (but still as early in beam as possible). NOTE: even this placement is not settled — upstream of M4/M5 won't happen, and before-M6 remains a real open question. Rotation as fast as possible; achromatic design; retardance 180° ± 1° across J–K
Beam Splitter
Wollaston prism in pupil plane (NIR) Dual-beam simultaneous I⊥ and I∥ on same detector or adjacent detectors May be several beam-splitters at the back-ends of different arms; split states must be eigenvectors of the optical train between modulator and beam-splitter
Derotator
Avoid angles near 45° operationally Optimised silver or dielectric coatings; full Mueller matrix model per rotation state Angles logged in all science and calibration data headers Possible HWPs immediately before/after the K-mirror to manage polarisation state entering and leaving it
Calibration Pipeline
Mueller matrix model of full ELT+PCS optical path Internal source calibrations + unpolarised standard stars IRDAP-style public pipeline from first light Polarisation-aberration correction algorithm included
Science Channels
NIR (J/H/K bands): dual-beam Wollaston imaging polarimeter Visible (600–900 nm): dual-beam HWP-modulated imager Secondary mode: IFS spectropolarimetry for companion characterisation
Diagnostic
Report Uphi images in all science outputs Internal calibration source for routine checks LLOWFS for post-coronagraphic wavefront monitoring
Notes: SPEAKER NOTES – Slide 12: Proposed architecture
This is the conceptual architecture for PCS polarimetry that the trade study points to.
Starting from top: an HWP as the modulator, placed as far upstream as possible in the PCS beam, after M5/M6 if the telescope design allows, otherwise before M5. It rotates at ~2 Hz to cycle between Q+ and Q- states. The achromatic specification is critical, the IRDIS lesson shows that a few percent error in HWP retardance can reduce the polarimetric efficiency to near zero in some bands.
For beam splitting: a Wollaston prism in the pupil plane, validated by both IRDIS and GPI. This gives clean orthogonal beam separation without chromatic issues.
Derotator: we need to work within ELT's existing design, but operationally we should avoid the 45-degree position. The coatings should be specified to minimise retardance variation. And every rotation state must be logged.
The calibration pipeline should be specified as a deliverable at Phase A: a full Mueller matrix model, internal calibration source, unpolarised standards, IRDAP-style public pipeline, and beam-shift correction.
For science channels: NIR plus visible, with an optional IFS mode for companion spectropolarimetry.
And as a diagnostic: always compute and report Uphi images. If Uphi is non-zero, something is wrong with the calibration. This has been the key diagnostic for SPHERE and it will be equally important for PCS.

## Slide 16
Next Steps
PCS Polarimetry — immediate priorities
Near term (1–3 months)
Finalise WP3.4 Polarimetry Management Plan and circulate for review
Write detailed specification for HWP: retardance tolerance, wavelength range, position in beam
Begin optical path analysis: model IP contributions from each ELT mirror (M1–M5) using Mueller calculus
Review Anche et al. follow-up paper (when published) for ELT polarisation aberration simulations with segment errors
Identify and contact Rob van Holstein (van Holstein et al.) regarding IRDAP applicability to ELT/PCS
Medium term (3–6 months)
Deliver Mueller matrix sensitivity analysis: what level of calibration is needed to reach 10⁻³ IP correction?
Evaluate beam-splitting options in detail: Wollaston vs. cube beam splitter vs. polarisation grating + QWP
Design observing strategy: HWP cycle angles, calibrator star cadence, standard star fields
Contribute to PCS Phase A optical design with polarimetry requirements formalised
Assess derotator options and work with instrument team on coating specifications
Longer term (> 6 months)
Develop prototype polarimetric data reduction pipeline (proof of concept)
Commission test observations with existing instrument (SPHERE/IRDIS) to validate calibration methodology
Prepare Phase A deliverables: polarimetry requirements, optical design inputs, calibration plan
Publish polarimetric design analysis as a contribution to PCS CDR documentation
ELT PCS | Polarimetry Work Update | June 2026
16
Begin science-case analysis; start in-house IP/aberration simulations (AI-assisted coding helps)
Notes: SPEAKER NOTES – Slide 13: Next steps
In the near term are the immediate priorities over the next one to three months. The main task is completing and circulating the WP3.4 management plan. Then starting the optical path analysis: modelling what IP contribution we expect from each mirror in the ELT optical path. This feeds into the specification for the HWP.
I also want to connect with Rob van Holstein at ESO, who is the main author of the IRDIS calibration papers. He can probably advise on what's needed to adapt IRDAP for ELT.
Medium term: the key deliverable is the Mueller matrix sensitivity analysis, essentially a calculation of what calibration accuracy we need, and what tolerances we can accept on each optical component. Also in this period: evaluating the different beam-splitting options and starting to formalise the requirements for the Phase A optical design.
Longer term: prototype pipeline, and if possible some test observations with SPHERE to validate the methodology before we apply it to ELT.

## Slide 17
Conclusions
PDI is the gold standard for all Polarimetric Imaging, PCS must include it from Day 1
The single most important lesson: dual-beam simultaneous design + HWP modulator (as fast as possible) + Mueller matrix calibration
Derotator and inclined Nasmyth mirrors are the primary systematic risks: must be modelled per exposure
ELT F/0.9 makes polarisation aberrations ~20× worse than SPHERE, requires new calibration strategies
To reach 10⁻³ IP correction for rocky planet science: full Mueller matrix model + upstream HWP is the path forward
Trade study table now in WP3.4 document (Section 7), full technical details available
Actual (spectro)polarimetry — Q/I, U/I — delivers the real science: microphysical characterisation of exoplanets and disks, beyond what PDI alone gives
Thank you — questions welcome
Notes: SPEAKER NOTES – Slide 14: Conclusions
Let me summarise the key points.
First, PDI is not optional for PCS, it should be considered the primary technique for disk science and it must be designed in from the beginning.
Second, and this is the big take-away from the entire trade study, the three pillars of a successful polarimetric design are: dual-beam simultaneous detection with the same or adjacent pixels; an HWP modulator at the right place in the beam; and a full Mueller matrix calibration model.
Third: the derotator and the Nasmyth mirrors are the primary systematic risks. We need to understand and model these from Phase A.
Fourth: ELT is genuinely harder than VLT. The faster focal ratio and the segmented primary both introduce challenges that don't exist for SPHERE. We need to engage with the Anche et al. programme and think carefully about what compensation strategies are available.
Fifth: for rocky planet science, the most ambitious case, correcting IP to 10⁻³ is the fundamental requirement. The path to get there is clear from the trade study: upstream HWP plus full Mueller matrix model.
Finally: all the detailed information from the trade study, all 7 instruments, what worked and what didn't, is now in Section 7 of the WP3.4 Polarimetry Management Plan. Please take a look.
Thank you. I'm happy to take questions.

## Slide 18
Discussion Topics for Team Meeting
Poster, slides, and WP3.4 document are circulated in advance — items to confirm with the group
ELT PCS | Polarimetry Work Update | June 2026
18
These materials go out ahead of time — please come prepared to discuss the open items below:
1
Poster
• Does the polarimetry group have any additions or changes to the poster?
2
Slides
• Trade-off: K-mirror avoidance strategy vs. VAMPIRES' QWP approach — VAMPIRES has added a QWP to recover polarimetric efficiency lost to derotation (unpublished)
• Calibration source strategy, list of multiple calibration sources vs. exactly one unpolarised + one polarised reference source
3
WP3.4 Document
• Technical review of poster content — accuracy, completeness, appropriate detail for a conference audience
• SPIE acknowledgements, funding statements, and grant numbers
• Beam-splitting architecture: feature in the poster, or reserve for oral presentation?
• Additional science cases or open questions to highlight for the SPIE community